
Terms and Conditions of Supply

Project Name: Genomic, spatial, and epidemiological analysis to inform targeted TB interventions in Moldova
Purchaser: Center for Health Policies and Studies 

1. Prices and Schedules for Supply

	Nr.
	Name of the product
	Measure unit
	Quantity
	Unit Price
	Total Price    
	Delivery Time

	1
	System for the preparation of media (+85ºC)
	unit
	2
	
	
	

	2
	Microtubes with detachable threaded cap for microcentrifuge
	piece
	5000
	
	
	

	3
	Microtubes with detachable threaded cap for microcentrifuge
	piece
	5000
	
	
	

	4
	Cryogene tubes 1,5 ml
	unit
	5000
	
	
	

	5
	Cryogene tubes 1,0 ml
	unit
	5000
	
	
	

	6
	CryoCaps with O-ring, neutro
	piece
	5000
	
	
	

	7
	Glass tubes 12cm
	piece
	1000
	
	
	

	8
	Glass tubes 16cm
	piece
	5000
	
	
	

	9
	Liquid thermometer
	piece
	10
	
	
	

	10
	Minimum and maximum digital thermometer -minus -30- plus + 50
	piece
	10
	
	
	

	11
	Minimum and maximum digital thermometer - minus -80- + 50
	piece
	10
	
	
	

	12
	Maximum liquid thermometer
	piece
	5
	
	
	

	13
	Automatic pipettes with adjustable volume 1000-5000 mcl
	piece
	2
	
	
	

	14
	Automatic pipettes with adjustable volume 100-1000 mcl
	piece
	2
	
	
	

	15
	Automatic pipettes with adjustable volume 1-10mcl
	piece
	2
	
	
	

	16
	Automatic pipettes with adjustable volume 20-200 mcl
	piece
	2
	
	
	

	17
	Aluminum foil
	piece
	5
	
	
	

	18
	Gloves (100 pieces per box), size S
	piece
	1,000
	
	
	

	19
	Gloves (100 pieces per box), size M 
	piece
	2000
	
	
	

	20
	Gloves (100 pieces per box), size L
	piece
	500
	
	
	

	21
	Fit Test Kit FT-30 Bitrex
	kit
	5
	
	
	


Note: In case of discrepancy between unit price and total derived from unit price, the correction will be done as provided in Paragraph 7 (iii) of the Invitation to quote

2. Fixed Price:  The prices indicated above are firm and fixed and not subject to any adjustment during contract performance.

3. The Purchaser reserves the right at the time of contract finalization to increase or decrease by up to 15% the quantity of goods and services originally specified without any change in unit prices as other terms and conditions.

4. Country of Origin: The goods offered should have their origin in World Bank member countries, and you will be required to furnish a certificate of origin for each item.

5. Delivery Schedule: The delivery should be not later than December 1, 2020. The Supplier shall be the sole responsible for failure to comply with the obligations he undertook in the agreement, the PAS Center is entitled to deduct the equivalent of 0,1% per day of delay out of the total value of the agreement but not more than 5% of the contract value. This amount will replace penalties as such.
6. Applicable Law: The Contract shall be interpreted in accordance with the laws of the Purchaser's country.

7. Resolution of Disputes: The Purchaser and the Supplier shall make every effort to resolve amicably by direct informal negotiation any disagreement or dispute between them under or in connection with the Contract. In the case of a dispute between the Purchaser and the Supplier, the dispute shall be settled in accordance with the country procedures.

8. Delivery and Documents: Upon shipment, the Supplier shall notify the Purchaser the full details of shipment, including contract number, description of goods, quantity, etc. The Supplier shall submit the following documents to the Purchaser:
(i) delivery document duly filled in, signed and sealed.

(ii) manufacturer's or supplier's warranty certificate (if applicable);

(iii) certificate of origin;

(iv) certificate of quality.

9. Payment: Invoices will be paid in 3 (three) working days in an amount of 100% against shipping documents and signature of acceptance certificate.
10. Warranty: Goods offered should be covered by manufacturer’s warranty.
11. Packaging and Marking Instructions: The Supplier shall provide standard packing of the Goods as required preventing their damage or deterioration during transit to their final destination, as indicated in the Contract.
12. Defects: All defects will be corrected by the Supplier without any cost to the Purchaser within 30 days from the date of notice by Purchaser.
13. Force-Majeure:  The supplier shall not be liable for penalties or termination for default if and to the extent that its delay in performance or other failure to perform its obligations under the Contract is the result of an event of Force-Majeure.

For purposes of this clause, “Force-Majeure” means an events beyond the control of the Supplier and not involving the Supplier’s fault or negligence and not foreseeable. Such events may include, but not restricted to, act of Purchaser in its sovereign capacity, wars or revolutions, fires, floods, epidemics, quarantine restrictions, and freight embargoes.

If a Force-Majeure situation arises, the Supplier shall promptly notify the purchaser in writing of such condition and the cause thereof. Unless otherwise directed by the Purchaser in writing, the Supplier shall continue to perform its obligations under the Contract as far as is reasonably practical, and shall seek all reasonable alternative means for performance not prevented by Force-Majeure event. 

14. Required Technical Specifications

	Nr.
	Name of the product
	Specifications and Requirements
	Specification proposed by the Suppliers

	1
	System for the preparation of media (+85ºC)
	 Media preparation system with tube holders. Temperature range (metric) + 30- + 95ºC Digital temperature settings Digital time setting1 min - 96 hours Temperature deviation inside with the set no more than 2% Power supply-220-230V. Number of seats in the stand- 180 -200. Tube dimensions (length x diameter) 170mm x 18mm. Temperature stability- ± 0.1 ° C
	

	2
	Microtubes with detachable threaded cap for microcentrifuge
	Polypropylene conical tubes with detachable threaded lid and O-ring, volume 1.5 ml, D = 11 mm. The sets will be delivered in original packaging, secured, marked and labeled by the manufacturer, without prepackaging; (without skirt) + caps
	

	3
	Microtubes with detachable threaded cap for microcentrifuge
	Polypropylene tubes with detachable threaded lid and O-ring, (V = 2.0 ml, diameter 11 mm, with skirt) + caps. The sets will be delivered in original packaging, secured, marked and labeled by the manufacturer, without prepackaging;
	

	4
	Cryogene tubes 1,5 ml
	Cryogen tubes 1.5-1.8 ml High density, transparent polypropylene with matte area for marking on the outer wall of the tube. Sterile, threaded lid, non-fixed tube, with O-ring, resistant to solvents and aggressive solutions, mechanical actions (vertexing), and high temperatures. Conical tube base (with skirt)
	

	5
	Cryogene tubes 1,0 ml
	Cryogen tubes 1.0 ml High density, transparent polypropylene with a matte area for marking on the outer wall of the tube. Sterile, conical, with threaded cover, not fixed to the tube, with O-ring, resistant to aggressive solvents and solutions, mechanical actions (vortexing), and high temperatures. Conical tube base (without skirt)
	

	6
	CryoCaps with O-ring, neutro
	Threaded outer caps for microtubes - packaging - in bags, compatible with 1.5 and 2 ml volume tubes
	

	7
	Glass tubes 12cm
	Glass tubes, transparent, with round bottom and plastic threaded lid. Autoclavable. Length 12 cm. Internal diameter 13mm (same along the entire length of the tube), external diameter not more than 20mm.
	

	8
	Glass tubes 16cm
	Glass tubes, transparent, with flat bottom and plastic screw cap. Autoclavable. Length 16 cm. Diameter 18mm.
	

	9
	Liquid thermometer
	Monitored temperature range -50 + 10 ° C, the value of a division 0.1 ° C. The sets will be delivered in original packaging, marked and labeled by the manufacturer, with metrological certificate
	

	10
	Minimum and maximum digital thermometer -minus -30- plus + 50
	Minimum and maximum digital thermometer; Temperature range -30… + 50˚C / resolution 0.1˚C; Internal memory for minimum and maximum; with alarm sensor. Can be used in rooms, refrigerators, freezers, incubators. To be delivered with a support for attachment on the refrigerated display case or on the incubator display case. To be delivered accompanied by the declaration of conformity from the manufacturer; To be delivered with the European standard EN13485 regarding the temperature control of the products (transport and storage); with Calibration / Calibration Certificate
	

	11
	Minimum and maximum digital thermometer - minus -80- + 50
	Minimum and maximum digital thermometer; Temperature range -80˚C- + 50˚C resolution 0.1˚C; Internal memory for minimum and maximum; with alarm sensor.  To be delivered with the declaration of conformity from the manufacturer; To be delivered with the European standard EN13485 regarding the temperature control of the products (transport and storage); Mandatory to be delivered with a Calibration / Calibration Certificate
	

	12
	Maximum liquid thermometer
	Monitored temperature range -10 + 150 ° C, the value of a division 0.1 ° C. The sets will be delivered in original packaging, marked and labeled by the manufacturer, mandatory with metrological certificate
	

	13
	Automatic pipettes with adjustable volume 1000-5000 mcl
	Micropipette with variable volume 1.0-5ml. The sets will be delivered in original packaging, marked and labeled by the manufacturer, with quality certificate or metrological certificate; The instructions for use of the kits should contain the performance and quality characteristics.
	

	14
	Automatic pipettes with adjustable volume 100-1000 mcl
	Micropipette with variable volume 100-1000mcl. The sets will be delivered in original packaging, marked and labeled by the manufacturer, with quality certificate or metrological certificate; The instructions for use of the kits should contain the performance and quality characteristics.
	

	15
	Automatic pipettes with adjustable volume 1-10mcl
	Micropipette with variable volume 1-10mcl. The sets will be delivered in original packaging, marked and labeled by the manufacturer, with quality certificate or metrological certificate; The instructions for use of the kits should contain the performance and quality characteristics.
	

	16
	Automatic pipettes with adjustable volume 20-200 mcl
	Micropipette with variable volume 20-200mcl. The sets will be delivered in original packaging, marked and labeled by the manufacturer, with quality certificate or metrological certificate; The instructions for use of the kits should contain the performance and quality characteristics.
	

	17
	Aluminum foil
	Dimensions 0.25-0.5 mm x 1000 x 25m. The sets will be delivered in original packaging, secured, marked and labeled by the manufacturer, without prepackaging; The instructions for use of the kits should contain the performance and quality characteristics.
	

	18
	Gloves (100 pieces per box), size S
	Disposable nitrile, powder-free gloves with a textured surface, high flexibility, tactility, slip resistance, tearing and chemical resistance. 
	

	19
	Gloves (100 pieces per box), size M 
	Disposable nitrile, powder-free gloves with a textured surface, high flexibility, tactility, slip resistance, tearing and chemical resistance.
	

	20
	Gloves (100 pieces per box), size L
	Disposable nitrile, powder-free gloves with a textured surface, high flexibility, tactility, slip resistance, tearing and chemical resistance.
	

	21
	Fit Test Kit FT-30 Bitrex
	Fit Test Kit FT-30 Bitrex, for respirators testing

The 3M Fit Test Kit used with any particulate respirator or gas and vapor respirator with a particulate filter to test the seal from face piece to face seal. Qualitative Fit Test Apparatus. Included: Sensitivity Solution (FT-31), Fit Test Solution (FT-32), 2 nebulizers (FT-13), 1 test hood (FT-14), 1 collar assembly (FT-15), and 2 sets of replacement nebulizer inserts. 
	



Supplier confirms compliance with above specifications {In case of deviations supplier to list all such deviations}


15. Failure to Perform: The Purchaser may cancel the Agreement if the Supplier fails to deliver the Goods, in accordance with the above terms and conditions, in spite of a 21-day notice given by the Purchaser, without incurring any liability to the Supplier. 

NAME OF SUPPLIER________________________________________________


Authorized Signature________________________________________


Date: 
PAGE  
2

